## Slide 1: Hemophilia Guidelines for All
A new ambition of the World Federation of Hemophilia

## Slide 2: WFH Guidelines for the Management of Hemophilia
All recommendations are consensus based.
Srivastava A et al. Haemophilia. 2020;26(Suppl 6):1–158.

## Slide 3: Chapter 3: Laboratory Diagnosis and Monitoring
Steven Kitchen PhD
Clinical Scientist
Department of Coagulation, Royal Hallamshire Hospital
Sheffield, United Kingdom

## Slide 4: Disclosure: Steven Kitchen
Consultant/Speaker fees
Novo Nordisk, Roche, Sobi ,Takeda, Werfen in the last 2 years.

## Slide 5: Authors
Steve Kitchen
Francisco de Paula Careta (PWH)
Silmara A. de Lima Montalvão
Emna Gouider
Radoslaw Kaczmarek (PWH)
Claude T. Tagny
Pierre Toulon
Glenn F. Pierce
Alok Srivastava
PWH, person with hemophilia

## Slide 6: Introduction
An accurate diagnosis depends on the laboratory following strict protocols and procedures including:
  Expertise in coagulation laboratory testing
  Correct equipment and reagents
  Quality Assurance
Diagnosis of hemophilia is based on:
  Understanding the clinical features of hemophilia and the appropriateness of the clinical diagnosis
  Screening tests: PT and APTT or platelet function tests
  Confirming the diagnosis by factor assays and other investigations
APTT, activated partial thromboplastin time; PT, prothrombin time.
Different bleeding disorders may have very similar symptoms; therefore, a correct diagnosis is essential to ensure that a patient receives the appropriate treatment.

## Slide 7: Coagulation laboratory testing
Recommendation 3.1.1
The WFH recommends that testing for diagnosis and monitoring of hemophilia must be carried out by staff with knowledge and experience in coagulation laboratory testing using equipment and reagents that have been validated for this specific purpose.
REMARK: Details of laboratory tests for the diagnosis and monitoring of hemophilia are described in the WFH laboratory manual.
The lab manual is provided to help laboratories to justify sufficient budget for proper set-up with administration
  Manual available on the WFH eLearning Platform
https://elearning.wfh.org/resource/diagnosis-of-hemophilia-and-other-bleeding-disorders-a-laboratory-manual/

## Slide 8: Coagulation laboratory testing
Recommendation 3.2.4 The WFH recommends that blood samples for determination of PT, APTT, or FVIII/FIX activity be collected in citrate tubes containing 0.105-0.109M (around 3.2%) aqueous trisodium citrate dihydrate, capped during processing, and kept at 18-25°C during transport and storage. Blood samples should be centrifuged at ambient temperature for a minimum of 1700 g for at least 10 minutes, and either be analyzed within 8 hours of collection (4 hours for FVIII:C) or stored deep frozen at −35°C or lower.
APTT, activated partial thromboplastin time; PT, prothrombin time.
REMARK: Storage of citrated whole blood samples at 2-8°C should be avoided as this may result in loss of FVIII activity.
REMARK: Platelet poor plasma (PPP) samples can be stored at −35°C for up to 3 months and at −70°C for up to 6 months prior to determination of FVIII/FIX activity. Storage of PPP at −20°C is usually inadequate. Freezers with auto-defrost should not be used to store PPP prior to determination of PT, APTT, or FVIII/FIX activity.

## Slide 9: Coagulation laboratory testing: Diagnostic
PT and APTT may be used to screen a patient suspected having a bleeding disorder
The sensitivity of both PT and APTT tests to factor deficiencies are influenced by the type of reagents used to perform the test
Recommendation 3.2.8 For laboratory investigation of patients being assessed due to clinical suspicion of hemophilia, the WFH recommends that an APTT result within the reference range not be used to rule out the presence of mild hemophilia A or B.
APTT, activated partial thromboplastin time; PT, prothrombin time.
REMARK: In some cases of mild hemophilia A or B, APTT may be within the normal range.

## Slide 10: Coagulation laboratory testing: Diagnostic
Recommendation 3.2.10 For laboratory investigation of patients being assessed due to clinical suspicion of hemophilia A, the WFH recommends the use of both the one-stage FVIII assay and the chromogenic FVIII:C assay in the initial diagnostic workup.
Some genetically confirmed cases have normal activity in one stage assays and reduced activity in chromogenic; the reverse may also occur
REMARK: Both assays should be performed even if the result of one of the two assays shows FVIII activity within the normal range.

## Slide 11: Coagulation laboratory testing: Assays
There are numerous published assay studies comparing results in samples containing CFCs including EHL FVIII and FIX concentrates showing important differences between results obtained with different assay methods
Assays that give results that differ by more than 25-30% from the labelled potency of the concentrate vial are best avoided or should not be used without taking account of such differences
CFC, clotting factor concentrate; EHL, extended half life.

## Slide 12: Coagulation laboratory testing: FVIII/FIX Assays
Recommendation 3.2.18 For monitoring replacement therapy with FVIII or FIX concentrates, the WFH recommends that laboratories use a FVIII/FIX assay that has been validated for use with the specific concentrate used for treatment.
REMARK: This recommendation is particularly important for modified molecular forms of FVIII and FIX.

## Slide 13: FVIII laboratory monitoring: Elocta®/Eloctate®
Recommendation 3.2.21 For monitoring replacement therapy with efmoroctocog alfa (recombinant FVIII fused with human immunoglobulin G1 [rFVIIIFc]; Elocta®/Eloctate®), the WFH recommends use of a one-stage or chromogenic FVIII assay calibrated with a plasma standard traceable to a WHO international standard.
Neither one-stage or chromogenic assays that have been studied have been shown to be unsuitable

## Slide 14: FVIII laboratory monitoring: Esperoct®
Recommendation 3.2.22 For monitoring replacement therapy with turoctocog alfa pegol (recombinant B-domain-truncated FVIII with a site-specific 40-kDa polyethylene glycol group [N8-GP]; Esperoct®), the WFH recommends use of a chromogenic FVIII assay or APTT-based one-stage FVIII with validated reagents, including some ellagic acid activator reagents (Actin®, Actin® FS, SynthAFax™, DG Synth™) and some silica activator reagents (Pathromtin® SL, SynthASil™), calibrated with a plasma standard traceable to a WHO international standard.
APTT, activated partial thromboplastin time; kDA, kilodalton.
REMARK: One-stage FVIII assays with APTT-SP™, STA®-PTT Automate, or TriniCLOT™ APTT HS reagents significantly underestimate true FVIII activity of N8-GP and should not be used.

## Slide 15: FVIII laboratory monitoring: Jivi®
Recommendation 3.2.23 For monitoring replacement therapy with damoctocog alfa pegol (recombinant B-domain-deleted FVIII with a site-specific 60 kDa branched polyethylene glycol group [BDD-rFVIII]; Jivi®), the WFH recommends use of a chromogenic FVIII assay or APTT-based one-stage FVIII assay with validated reagents, including the ellagic acid activator reagent Actin® FSL and some silica activator reagents (Pathromtin® SL, SynthASil™), calibrated with a plasma standard traceable to a WHO international standard.
REMARK: One-stage FVIII assays with the ellagic acid activator reagent Actin® FS or the kaolin activator reagent C.K. Prest® significantly overestimate true FVIII activity and should not be used. One-stage FVIII assays with APTT-SP™ and STA®-PTT Automate reagents significantly underestimate true FVIII activity and should not be used.
APTT, activated partial thromboplastin time; kDA, kilodalton.

## Slide 16: FVIII laboratory monitoring: Adynovate® /Adynovi®
Recommendation 3.2.24
For monitoring replacement therapy with rurioctocog alfa pegol (full-length recombinant FVIII with non-site-specific 20-kDa polyethylene glycol; Adynovate®/Adynovi®), the WFH advises that more laboratory assay studies are required to inform recommendations about laboratory monitoring.
There is no published evidence that any one-stage assay is unsuitable
However, there is conflicting published data for chromogenic assays
REMARK: There are conflicting findings in the literature assessing the use of one-stage and chromogenic FVIII assays in samples containing rurioctocog alfa pegol.
kDA, kilodalton.

## Slide 17: FVIII laboratory monitoring: Afstyla®
Recommendation 3.2.25
For monitoring replacement therapy with lonoctocog alfa (single-chain recombinant FVIII [rVIII-SingleChain]; Afstyla®), the WFH recommends use of a chromogenic FVIII assay calibrated with a plasma standard traceable to a WHO international standard.
REMARK: The summary of product characteristics recommends chromogenic assays. It also states that the one-stage FVIII assay result underestimates the FVIII activity level by approximately 45% compared to the chromogenic assay result and suggests that if the one-stage assay is used, the result should be multiplied by a factor of 2.

## Slide 18: FIX laboratory monitoring: Alprolix®
Recommendation 3.2.28
For monitoring replacement therapy with eftrenonacog alfa (recombinant FIX fused with human immunoglobulin G1 [rFIXFc]; Alprolix®), the WFH recommends use of a chromogenic FIX assay or APTT-based one-stage FIX assay with validated reagents, including some ellagic acid activator reagents (Actin®, Actin® FS, Actin® FSL), some silica activator reagents (Pathromtin® SL, SynthASil™), and a polyphenol activator reagent (Cephascreen®), calibrated with a plasma standard traceable to a WHO international standard.
APTT, activated partial thromboplastin time.
REMARK: One-stage FIX assays with STA®-PTT Automate or kaolin activator (C.K. Prest®) reagents significantly underestimate true rFIXFc (Alprolix®) activity and should not be used.

## Slide 19: FIX laboratory monitoring: Idelvion®
Recommendation 3.2.29
For monitoring replacement therapy with albutrepenonacog alfa (recombinant FIX fused with recombinant human albumin [rFIX-FP]; Idelvion®), the WFH recommends use of an APTT-based one-stage FIX assay with validated reagents, including some silica activator reagents (Pathromtin® SL, SynthASil™), calibrated with a plasma standard traceable to a WHO international standard.
REMARK: One-stage FIX assays with the ellagic acid activator reagent Actin® FS or the kaolin activator reagent C.K. Prest® significantly underestimate true rFIX-FP (Idelvion®) activity and should not be used. One-stage assays with the ellagic acid activator SynthAFax™ reagent or chromogenic FIX assays significantly overestimate true rFIX-FP (Idelvion®) activity and should not be used.
APTT, activated partial thromboplastin time.

## Slide 20: FIX laboratory monitoring: Refixia®/Rebinyn®
Recommendation 3.2.30
For monitoring replacement therapy with nonacog beta pegol (recombinant FIX with a 40-kDa polyethylene glycol moiety [N9-GP]; Refixia®/Rebinyn®), the WFH recommends use of a chromogenic FIX assay or APTT-based one-stage FIX assay with validated reagents, including the ellagic acid activator reagent SynthAFax™ or the polyphenol activator Cephascreen®, calibrated with a plasma standard traceable to a WHO international standard.
REMARK: Most one-stage FIX assays significantly overestimate or underestimate true FIX activity of N9-GP and should not be used. One-stage assays using the ellagic acid activator reagent SynthAFax™ or the polyphenol activator reagent Cephascreen®, are suitable for monitoring therapy with N9-GP.
APTT, activated partial thromboplastin time; kDA, kilodalton.

## Slide 21: FVIII laboratory monitoring: Emicizumab
An engineered bispecific antibody that binds both human FIX/FIXa and FX/FXa and a FVIII mimetic
The APTT is shortened by emicizumab to within or below the ref range, irrespective of reagent
Emicizumab significantly interferes in chromogenic FVIII assays utilizing human FIXa and FX, but not those using bovine FIXa and FX
Emicizumab can be measured and reported in μg/mL using a modified one-stage assay with higher test sample dilution and calibrated with drug-specific calibrators
APTT, activated partial thromboplastin time.
Phospholipid membrane
Bispecific antibody
HC
LC
FX
HC
LC
FIXa

## Slide 22: FVIII laboratory monitoring: Emicizumab
Recommendation 3.2.31
For patients receiving emicizumab in whom confirmation of expected emicizumab levels is required, the WFH recommends use of a modified one-stage assay including an additional pre-dilution step of test plasma and assay calibration with specific emicizumab calibrators.
APTT, activated partial thromboplastin time.
REMARK : Even at subtherapeutic levels of emicizumab, APTT may be normal or subnormal in patients with severe hemophilia A with or without inhibitors.

## Slide 23: FVIII laboratory monitoring: Emicizumab
Recommendation 3.2.32
For determination of FVIII activity in patients with hemophilia A receiving emicizumab, the WFH recommends use of a chromogenic FVIII assay containing bovine FX.
REMARK : At therapeutic levels, emicizumab affects any chromogenic FVIII assay containing FX of human origin. Emicizumab may also affect chromogenic FVIII assays containing FIXa of human and FX of bovine origin but only at emicizumab levels higher than those expected in patients receiving recommended doses.

## Slide 24: FVIII laboratory monitoring: Emicizumab
REMARK: Validated anti-drug antibody assays may also be used for this purpose, if available.
Recommendation 3.2.33 For determination of FVIII inhibitor levels in patients receiving emicizumab, the WFH recommends use of a chromogenic FVIII assay containing bovine FX.
Recommendation 3.2.34 For patients with a suspected neutralizing anti-emicizumab antibody, the WFH recommends measuring emicizumab levels using a modified one-stage assay including an additional pre-dilution step of test plasma and assay calibration with specific emicizumab calibrators.

## Slide 25: Laboratory monitoring: Inhibitors
The most frequently encountered functional inhibitors of hemostasis are lupus anticoagulants
Most FVIII inhibitors that develop secondary to replacement therapy in patients with hemophilia A show a characteristic pattern: the APTT of a patient/PNP mixture is intermediate
If there is suspicion of lupus anticoagulant or if the sample contains therapeutic anticoagulants, it may be useful to confirm inhibitor presence using a chromogenic assay to measure residual factor activity
APTT, activated partial thromboplastin time; PNP, pooled normal plasma.

## Slide 26: Laboratory monitoring: Inhibitors
Recommendation 3.2.37 For quantification of anti-FVIII inhibitors, the WFH recommends that the Nijmegen-Bethesda assay be used.
REMARK: Bethesda assays detect neutralizing antibodies. A small proportion of anti-FVIII antibodies are non-neutralizing, shorten the half-life of infused FVIII, and are not detected by Bethesda assays.
REMARK: The Nijmegen modification describes a specific method for buffering pooled normal plasma; other buffering methods may be suitable.

## Slide 27: Laboratory monitoring: Inhibitors
Buffered normal plasma as source of FVIII is mixed with:
  FVIII deficient plasma containing VWF (control mix)
  Patient sample (first destroy any FVIII by heat treatment)
Determine residual FVIII after 2 hours at 37oC in test compared to control mix
VWF, von Willebrand factor.

## Slide 28: Laboratory monitoring: Inhibitors
Recommendation 3.2.35
For determination of anti-FVIII inhibitors in a sample containing greater than 5 IU/dL FVIII activity, the WFH recommends that prior to testing, the sample be heated to 56°C for 30 minutes and centrifuged at ambient temperature for a minimum of 1700 g for at least 5 minutes.
REMARK: The quantification limit of the Nijmegen-Bethesda FVIII inhibitor assay is around 0.6 BU/mL.
REMARK: The Nijmegen-Bethesda FVIII inhibitor assay requires use of buffered pooled normal plasma as a source of FVIII, which is then mixed with an equal volume of FVIII-deficient plasma to prepare the control mixture.

## Slide 29: Laboratory monitoring: Quality assurance
Internal quality control is used to establish whether a series of techniques and procedures is being performed consistently over time
External quality assessment helps identify the degree of agreement between the local laboratory results and those obtained by other centres
  For a laboratory to attain a high level of testing reliability and to participate successfully in an external quality assessment program, the laboratory must have access to:
    Appropriate reagents and techniques
    Appropriate number of adequately trained staff

## Slide 30: Laboratory monitoring: Quality assurance
Recommendation 3.4.1 The WFH strongly recommends that coagulation laboratories implement quality assurance programs for all laboratory systems to ensure quality adherence and the reliability of laboratory blood testing procedures and reporting for the diagnosis and treatment of hemophilia.
Recommendation 3.4.3 The WFH strongly recommends that clinical laboratories routinely participate in external quality assessment for each assay used for the diagnosis and treatment of hemophilia.
REMARK: Participation in the WFH International External Quality Assessment Scheme (IEQAS) enables laboratories to improve and standardize laboratory testing for hemophilia.

## Slide 31: Conclusions
Accurate laboratory testing for diagnosis and monitoring of hemophilia requires:
Close collaboration between laboratory and clinical service
Access to sufficient appropriate reagents and equipment
Resources to follow the WFH guidance on laboratory diagnosis and monitoring
Laboratory staff with knowledge and expertise in coagulation to select and implement correct methods

## Slide 32: Thank you to the Hemophilia Alliance for their support in developing this presentation